RICS and Scanning FCS Analysis Plugins

Jay Unruh

Research Advisor

Stowers Institute for Medical Research

Note: please read the License Information file for details about licensing and use.  By using this software, you are agreeing to the enclosed license agreement.

Theory

RICS analysis was formally discovered by Digman et al in 2005 in Enrico Gratton’s laboratory with contribution from the Paul Wiseman group.  Nevertheless, the method is closely related to scanning FCS which was proposed by Berland et al in 1996, also in Enrico’s lab.  In fact, it is straightforward to derive the RICS equation from Berland’s 1996 equation.  The basic problem is as follows:

Consider an open volume with freely diffusing fluorescent particles.  If we were to collect a laser scanning confocal image of this volume, what we would see will depend on the rate at which the particles diffuse (related to the diffusion coefficient).  If the particles diffuse much faster than the scanning of the laser beam, we will see “streaks” in the image, approximately one line in height.  If the particles diffuse much slower than the scanning of the laser beam, we will see “blobs” in our image corresponding to each particle.  Between these extremes lies a continuum of diffusion rates.  So how do we quantitate this information?  The answer, of course, is to perform image correlation.  For fast moving particles, image correlation gives us something that looks like this:

The correlation function is given as follows:

G(ξ,ψ)=δI(x,y)δI(x+ξ,y+ψ)I2

So what does this mean?  As a rule of thumb, the correlation function gives us the average “shape” of our streaks.  In other words, the intensity at each point in the correlation function is proportional to the probability of the molecule still being in the confocal volume at distance (x,y) away from its original position and at time (t) from its original discovery.  The further away you go from the origin, the less likely the molecule is to be in focus.  By the next line that is acquired, a large amount of time has passed and the molecule will likely have diffused away.

So how do we analyze this data.  In general, we just need to calculate the spatio-temporal probability distribution for our molecule.  This distribution was written down long ago and is given by:

P(r)∝expr24Dτ

Luckily for us, this is just a Gaussian!  There is one problem though, if D = 0, this distribution is a point in the middle of our correlation function, not a blob, like we predicted.  This blob is a result of the diffraction of light in our microscope, and must be accounted for.  The derivation is too involved to go into here, but the general form of the equation is given as follows:

G(ξ,ψ)=G(τ)S(r)

G(τ)=γN1+4Dτw021+4Dτz02

S(r)=exp−r2w02+4Dτ

Here G(t) is the familiar single point autocorrelation function from traditional FCS measurements.  S(r) is a scanning term, very similar to the one given by Berland et al.  I have modified the term for single photon excitation.  For two photon excitation, the above equations change as follows:

G(τ)=γN1+8Dτw021+8Dτz02

S(r)=exp−2r2w02+8Dτ

The S term is now identical to Berland et al.  You might ask, why is our original correlation term independent of t while both G and S terms are dependent on t?  The answer is that r and t are dependent on one another.  For raster scanning (RICS), t is given as follows:

r2=(ξ+ψ)2

τ=ξτp+ψτl

Where tp is the time per pixel and tl is the time per line.  We now have all of the information we need to perform scanning FCS analysis or RICS analysis.

Performing RICS Analysis

The RICS experiment, in principle is very simple.  One typically collects on the order of 25 to 50 scanning confocal images with a pixel dwell time of around 10 us and a pixel size of around 50 nm.  Images are typically collected in power of 2 sizes (128, 256, 512) for easy calculation of the correlation function with the fast fourier transform.  The average correlation of the images is then calculated.  Finally the correlation function is fit to the above equations to obtain the diffusion coefficient, number of particles, and perhaps the beam waist (w0).  This method works well for photon counting images of diffusing dyes.

Perhaps the greatest challenge to performing RICS analysis is ensuring that all spatial fluctuations in the image are molecular fluctuations and not intensity changes due to cell boundaries or bright aggregates.  In order to obtain these conditions, one typically must detrend the data.  There are two established ways to accomplish this.  The first is by simply subtracting the average image from the image series.  Since molecules of interest are diffusing, they will not be subtracted out while larger scale structures will be subtracted.  Unfortunately, cells move as well.  This means that subtracting the average will not eliminate all non-molecular motion.  The solution is to subtract a moving average.  The period of the moving average should be chosen as long as possible in order to avoid subtracting molecular features, but short enough to eliminate cellular motion.  For fast moving molecules, a moving average period of 5 is typically sufficient.

Of course, once we subtract the average or the moving average, we are left with an average intensity of zero.  This means our correlation function goes to infinity!  It also means our correlation amplitude is way off.  The solution is to add back the average intensity of each image before calculating the correlation function.  Note that this does not perfectly correct our correlation amplitude.  In dark regions of the original image, too much intensity will be added back.  In bright regions of the image, too little intensity will be added back.  This can lead to fairly dramatic errors in the correlation amplitude as well as things that are calculated from it (like concentration and molecular brightness).  Therefore, it is wise to start with an image that is as uniform as possible so that these measurements are not strongly skewed.

One other caveat to RICS analysis is the use of analog instrumentation.  The vast majority of commercial instruments on the market are analog instruments rather than photon counting instruments.  One must be very careful when working with analog instruments.  For example Zeiss LSM 510 and LSM 700 confocals are notorious for having dominant ringing effects that make RICS impossible.  Olympus likes to call their instrument “photon counting” when in reality, they simply scale the intensity by a calibration so that each intensity is approximately equal to the number of detected photons.  All analog instruments that I have seen have at least some intrinsic horizontal correlation between pixels.  Therefore it is wise to omit the first few points of the correlation along the horizontal axis in the analysis.  (The first point should always be omitted to avoid shot noise).  All analog instruments have an offset—even the good ones.  That offset makes no difference in the shape of the correlation curve, but it does result in correlation amplitude.  Therefore the offset should always be subtracted before performing RICS analysis.

Using the RICS Plugins

There are several plugins enclosed for doing RICS analysis.  To install the plugins, just copy them to your plugins folder (c:\Program Files\ImageJ\plugins on windows) along with jalgs.jar, j3D.jar, and jguis.jar.  Then restart ImageJ and they should appear under the plugins menu.  I have chosen to allow the user to choose the exact method for analysis since users have different analysis methods depending on their sample conditions.

First the user needs to import the data into ImageJ.  This will vary depending on the microscope manufacturer, but the new versions of ImageJ have import libraries for most systems.  If not, feel free to try the LOCI Bio-Formats importer.  If you still have trouble, export the data from your microscope as a tif series and then import into ImageJ.  You will also need to know the pixel dwell time and the time between lines for your confocal system.  The fitting program also asks for the pixel size, but this is populated from the image itself if the microscope saves it in the image.  Also note that my plugins only allow 16-bit or floating point images.  If you are having trouble, just convert to floating point (32-bit) to be safe (Image-Type-32-bit).

Once the image is imported, you will need to make decisions about detrending.  If there are any non-random variations in intensity over the image, you will need to detrend the image.  If no cell motion occurs during the movie, you can get by with average subtraction.  Otherwise, you can try moving average subtraction (stack subtract moving average jru v1).  Either way, make sure and select maintain spatial average so you don’t end up with a zero image.

Once the image is detrended, you can calculate the correlation function with stack FFT ICS jru v1 or stack FFT ICCS jru v1.  Inspect the correlation image carefully to make sure you didn’t miss immobile structures.

Finally, fit the correlation image with fit RICS jru v1 or fit RICS line jru v1.  The line version only fits the horizontal or vertical center line of the correlation function.  Most of the items here are fairly self-explanatory.  The fix checkboxes appear under and to the left of each parameter.  This is an unfortunate result of the ImageJ generic dialog class.  Note carefully the Pts in x to fit, Pts in y to fit, and Pts in x to skip.  You will want to skip enough points to eliminate the ringing present from your microscope.  This can be easily identified by collecting data with room light or something else that shouldn’t have correlation.  All that is left is the ringing component.  You will want to fit enough points in x to fit the tail of the correlation function.  Make sure you don’t fit so many points that you weight the fit heavily by the noise that is present at large distances away from the correlation origin.  Basically, you want the correlation to go to zero within the fit window, but not more than that.  On the fitting dialog, if you press OK, the fit will take place and then the dialog will reappear to show the results.  This will continue to happen until you click Cancel.  Then the results of the fit will be written to the log window.  There is a Test Chi Squared box that allows for adjustment of parameters without fitting.

References:

Berland, K. M., P. T. C. So, Y. Chen, W. W. Mantulin, and E. Gratton.  Scanning Two-Photon Fluctuation Correlation Spectroscopy: Particle Counting Measurements for Detection of Molecular Aggregation.  1996, 71:410.

Digman, M. A., C. M. Brown, P. Sengupta, P. W. Wiseman, A. R. Horwitz, and E. Gratton.  Measuring Fast Dynamics in Solutions and Cells with a Laser Scanning Microscope.  2005, 89:1317.

Digman, M. A., P. W. Wiseman, A. R. Horwitz, and E. Gratton.  Detecting Protein Complexes in Living Cells from Laser Scanning Confocal Image Sequences by the Cross Correlation Raster Image Spectroscopy Method.  2009, 96:707.