analysis_autocorr and analysis_crosscorr plugins for ImageJ

Jay Unruh

Research Advisor

Stowers Institute for Medical Research

Note: please read the License Information file for details about licensing and use.  By using this software, you are agreeing to the enclosed license agreement.

Theory:

Please see our FCS and PCH webpages (http://research.stowers-institute.org/microscopy/external/Technology/FCS/index.htm) for details of FCS theory.  There are a few things to add on top of what is available there in relation to cross correlation.  The bleedthrough fraction is often denoted as follows:

β=IrIg

Where Ir is the intensity of the green species in the red channel (bleedthrough) and Ig is the intensity of the green species in the green channel.  Given this definition, we can define a minimum cross correlation amplitude for non-interacting species:

G(0)cc,min=G(0)gβIgIr

This value makes it possible to easily determine whether or not interaction is present in the current experiment.  If the cross correlation amplitude is above the minimum value, then interaction is present.

Experimental Guidelines:

Most of the pitfalls in doing FCS and FCCS are related to laser power, photobleaching, or cell motion for in vivo analysis.  If laser power is too high, triplet dynamics and focal volume saturation will dominate the analysis and prevent accurate determination of dynamics, concentration, and molecular brightness.  If significant photobleaching or cell motion occurs, much of the observed dynamics will be the slow dynamics of these processes and not the molecular dynamics desired.  Slow dynamics are especially disastrous for cross correlation analysis.  If both the red and the green channels experience bleaching or cell dynamics, they will be strongly correlated and it will appear as though a slow pool of interacting molecules exists.  The same is true for molecular brightness or concentration analysis (see the PCH website).  Slow dynamics here lead to much more dramatic fluctuations and therefore an overestimation of the molecular brightness (and an underestimation of the number of particles).

Running the Program:

To install the program, simply extract all of the files from Jay Plugins.zip to your ImageJ plugins folder (c:\Program Files\ImageJ\plugins\ on windows systems).  If you do not have ImageJ, it is freely downloadable from http://rsbweb.nih.gov/ij/download.html.  The plugins “analysis autocorr” and “analysis crosscorr” will then appear under the ImageJ plugins menu.  When you click on this menu item, a dialog will appear asking for the sampling frequency and the psf type (3D Gaussian or 2D Gaussian for membrane diffusion).  Then a file open dialog appears allowing for the selection of multiple confocor 3 raw files (hold down control and click multiple files).  There are a few other options for file import.  Short binary trajectory reads in a binary file with a list of unsigned 16-bit integers giving the number of photons per time bin.  PlotWindow trajectory works with trajectories generated by my simulation tools.  Note that other than for confocor 3 raw files, the program does not care about the sampling frequency of the histogram.  Segmented linear detrending is available for attempts at removing slow dynamics.  If this option is selected, all curves will be fit to a segmented line (the number of segments can be varied).  Then that line is subtracted from the curve and the average intensity is added back before the correlation is calculated.  Note that if too many segments are used and slow molecular fluctuations are present in the analysis, detrending could eliminate slow molecular fluctuations.  Also note that detrending is intended for correction of minor slow dynamics.  Curves with large scale dynamics should be eliminated rather than detrended.